|  | Document Title: |  | Revision Date: | 6/4/2026 |
| --- | --- | --- | --- | --- |
|  | Cannabinoid Potency Assay SOP |  | Version Number: | 1.1 |
|  |  |  | Approved By: | Brad Towey |
|  |  |  | Department/Category: | Laboratory |
|  | Document Type: | SOP | Page(s): | 1 of 12 |

# Purpose

- The purpose of this document is to provide instruction and to establish training standards regarding cannabinoids potency testing of cannabis flower and cannabis derived products using HPLC.

- Table 1. Assay Measurands

| - Assay Measurands | - Type | - Description |
| --- | --- | --- |
| Cannabinoids (µg/mL or mg/g) | Potency Assay | Cannabinoids are identified if their retention time corresponds to the retention time of any standard, QC, or CCV injection and quantified against an external calibration curve. |

# Scope

This SOP covers general assay setup cannabis samples, instrument setup, and calculations for chemists and laboratory personnel at AIT.

The scope of this analytical method is to instruct personnel on the execution of the HPLC method to quantitatively determine the cannabinoids present in cannabis plants and non-flower derived products. The samples are split into five different categories:

Plants (Flower): samples include any hemp or marijuana plant part, biomass, stems, stalks, leaves, seeds, as well as the dried flower portion of pre-rolls, etc.

Concentrates – Processed/extracted material from marijuana or hemp includes distillates, waxes, hash, isolates, emulsions, tinctures, lotions, etc.

Beverages – Cannabinoid infused soft drinks, teas, coffees, etc.

Edibles – Cannabinoid infused gummies, chocolates, candies, butter, etc.

Transdermal Patch – A concentrate matrix applied to a patch.

Cannabinoids in the scope of this SOP are:

          - Delta 9-tetrahydrocannabinol (d9-THC)

          - Delta 8-tetrahydrocannabinol (d8-THC)

          - Tetrahydrocannabinolic acid (THCA)

          - Tetrahydrocannabivarin (THCV)

          - Tetrahydrocannabivarinic acid (THCVA)

          - Cannabidiol (CBD)

          - Cannabidiolic acid (CBDA)

          - Cannabidivarin (CBDV)

          - Cannabigerol (CBG)

          - Cannabigerolic acid (CBGA)

          - Cannabinol (CBN)

          - Cannabichromene (CBC)

          - Cannabicyclol (CBL)

# Roles and Responsibilities

## Chemists

- AIT chemists are responsible for reading, understanding, maintaining, and performing the tasks in compliance with the expectations of this SOP. Personnel are also expected to report any problems with the execution of this procedure to the Senior Analytical Chemist and/or the COO.

# References

Table 2. Document References

Document Title

Cannabinoid Potency Assay Form

Sample Weighing Log

Sample Dilution Factor

Minnesota Office of Cannabis Management – Cannabis Technical Authority

# Definitions

Table 3. Defined terms

| Term | Definition |
| --- | --- |
| HPLC | High Performance Liquid Chromatography |
| CCV | Calibration Curve Verification |
| LCS | Laboratory Control Samples |
| ACN | Acetonitrile |
| IPA | 2-proponal, Isopropanol, or iso-propyl alcohol |
| PDA | Photodiode Array |

# Safety

## Ensure the following safety precautions are followed:

Lab coats, safety glasses, and gloves shall be worn while performing this procedure.

# Materials, Equipment, and Software

Table 4. Recorded Materials

| Identifier | Catalog Number |
| --- | --- |
| Shimadzu NexLeaf CBX for Potency 2.7 µm, 150 x 4.6 mm | 220-91525-70 |
| Shimadzu SH-EXP Guard Holder | 220-91525-73 |
| Shimadzu NexLeaf Guard 5 x 4.6 mm | 220-91525-72 |
| Water, LC-MS Grade (Honeywell) | LC365-4 |
| Formic Acid, LC-MS Grade (Fisher) | A117-50 |
| Acetonitrile, HPLC Grade (Concord Technology) | 8002LC4000 |
| Isopropanol, HPLC or LC-MS Grade (Sigma, Honeywell) | 34863 or LC323-4 |
| Methanol, HPLC or LC-MS Grade (Concord Technology, Honeywell) | 8001LC4000 or LC230-4 |
| Custom 13-Phytocannabinoid Mixture CRM - AIT (Cayman Chemical Company) | 9004931 |
| Olive Oil (Absolute Standards, Inc., Thermo Scientific) | 51340 or 416530250 |
| Cellulose (MP Biomedicals) | 191499 |
| QuEChERS 4 g MgSO4 + 1 g NaCl (Agilent, Restek) | 5982-6550 or 25847 |
| QuECHERS 6 g MgSO4 + 1.5 g NaCl (Sigma, Restek) | 55295-U or 28898 |
| 80:20 Acetonitrile:Water Solution (Prepared by AIT) | YYMMDD-8020 |

- Table 5. Non-recordable Materials

| Name | Item Number |
| --- | --- |
| Pipette Tips, various sizes | Various |
| 5, 15, and/or 50 mL tubes | Various |

- Table 6. Equipment

| Name | EQ No. |
| --- | --- |
| Shimadzu LC-2030C Plus; Software: LabSolutions Version 5.117 | Various |
| Analytical Balance | Various |
| Moisture Balance | Various |
| Dispenser and Pipettes (air & positive displacement, repeater) | Various |
| Vortexer | Various |
| Sonicator | Various |
| Centrifuge | Various |

- Table 7. Non-recordable equipment

| Name | EQ No. |
| --- | --- |
| 0.22 µm PTFE syringe filters | Various |
| 3 and/or 5 mL syringes | Various |

# Procedure

- The preparation of any of these solutions may be scaled as necessary if the ratios of each component(s) are maintained.

## Documentation and Sample Information

To ensure proper documentation of materials (Table 4), equipment (Table 6), and assay execution, use the “Cannabinoid Potency Assay Form.”

A unique Assay ID in the format of YYYYMMDD_CP is used to link the form’s multiple pages, excel worksheets, and data printouts.

- Where:

DD: Day

MM: Month

YYYY: Year

CP: Cannabinoid Potency

## Preparation of Mobile Phase A

Add 2.0 mL of formic acid to 2 L of water. Sonicate for 5 minutes to de-gas the solution.

Label the solution with the Solution ID in the format of YYYYMMDD_CP_MPA.

Expiration date is three months from the date of preparation.

Store at Room Temperature.

## Preparation of Mobile Phase B

Add 5.0 mL of formic acid to 5 L of acetonitrile. If desired, sonicate for 5 minutes to de-gas the solution.

Label the solution with the Solution ID in the format of YYYYMMDD_CP_MPB.

Expiration date is three months from the date of preparation.

Store at Room Temperature.

## Preparation of 80:20 ACN:H2O

Mix 200 mL of ACN with 50 mL of H2O.

Label the solution with the Solution ID in the format of YYYYMMDD_CP_8020.

Store at Room Temperature.

## Preparation of the Standard Curve and CCV Samples

Follow Table 8 for the preparation of the standard curve and CCV samples.

- Table 8. Preparation of the Standard Curve and CCV Samples.

| Sample ID | Spiking Solution (µL) | Amount of Spiking Solution (µL) | Volume ACN (µL) | Concentration (µg/mL) |
| --- | --- | --- | --- | --- |
| STD 6 | 13 Cannabinoid Mix Stock Solution (500 ppm) | 77.0 | 77.0 | 250 |
| STD 5 | STD 6 | 54.0 | 81.0 | 100 |
| STD 4 | STD 5 | 33.0 | 99.0 | 25 |
| STD 3 | STD 4 | 32.0 | 128.0 | 5 |
| STD 2 | STD 3 | 60.0 | 90.0 | 2 |
| STD 1 | STD 2 | 50.0 | 50.0 | 1 |
| CCV | 13 Cannabinoid Mix Stock Solution (500 ppm) | 200.0 | 4,800 | 20 |
| LOQ | 13 Cannabinoid Mix Stock Solution (500 ppm) | 10.0 | 4,990 | 1 |

Label each solution with the Solution ID in the format of YYYYMMDD_CP_(Sample ID).

Once the curve (STD 1-6) is run on the HPLC, the curve is stored on the instrument. A calibration curve is not needed for each run if Quality Control Requirements are met, and the drift of the instrument can be monitored with a CCV sample. It is recommended that a new curve should be prepared and run after preventative maintenance was performed, a different column was added, if the UV bulb was replaced, or any other major work performed on the instrument.

Store at -20 °C ± 10 °C when not in use.

The expiration date of this working solution is the same as the stock solution.

## Preparation of Laboratory Control Samples

LCSs should be prepared fresh each day if their respective sample type is to be assayed.

- Table 9. Preparation of LCSs.

| Sample ID | Sample Type | Matrix | Amount of Matrix | Amount of 12 Cannabinoid Mix Stock Solution |
| --- | --- | --- | --- | --- |
| PLC | Flower/Plant | Cellulose | 500 mg | 60.0 µL |
| CLC | Concentrate | Olive Oil | 38.0 µL | 12.0 µL |
| BLC | Beverage | Carbonated Water | 10 mL | 60.0 µL |
| ELC | Edibles | Gelatin Powder | 50.0 mg | 60.0 µL |

All LCSs have a final concentration of 3 µg/mL, except BLC at 1.5 µg/mL, after they have been processed following their respective sample type extraction.

## Preparation of Flower/Plant Samples

DO NOT trim or dry the plant samples.

Use a clean grinder/blender cup to grind the rest of the sample until they are visually homogenous. Transfer the ground sample to a labeled container.

Weigh 500.0 mg ± 10.0 mg of sample into a labeled, 50 mL tube.

Add 20.0 mL of ACN to the sample container.

Vortex sample at speed 10 for 5 minutes, then sonicate for 5 minutes.

Vortex sample at speed 10 for 1 minute, then sonicate for 5 minutes.

Centrifuge the samples at ≥3000 rpm for 5 minutes.

Attach a clean, 0.2 µm PTFE filter to a syringe and dispense contents into a new 5 mL tube (DF=40).

Dilute the samples with ACN (if needed). Common dilution factors include:

- Hemp Flower: 400x – 600x

- Marijuana Flower: 1000x – 1500x

- Proficiency Hemp Flower: 100x & 500x

## Preparation of Concentrate Samples

Weigh 50.0 mg ± 1.0 mg of sample into a labeled, 5 mL tube.

Add 0.950 mL of IPA to the sample container.

Vortex sample at speed 10 for 5 minutes, then sonicate for 5 minutes.

Vortex sample at speed 10 for 1 minute, then sonicate for 5 minutes.

Add 1.000 mL of ACN to the sample container.

Vortex the sample at speed 10 for 1 minute.

If necessary, syringe filter the supernatant using a 0.2 µm PTFE filter and dispense contents into a 5 mL tube (DF=40).

Dilute the samples with ACN (if needed). Common dilution factors include:

- Distillates, Hydrocarbon Wax, Live Resin, Live Rosin, RSO, etc.: 4000x

- Tinctures, Capsules, Capsule Solutions, etc.: 300x

- Topicals: 200x

- Emulsions: 40x – 1000x

- Isolates: 5000x

## Preparation of Beverage Samples

Open the beverage and pour out the contents in a suitably sized beaker.

Add a magnetic stir bar and stir for 10 minutes at 300-600 rpm.

Place beakers into a sonicator and sonicate for 10 minutes at ~35°C.

Transfer 10.0 mL of the beverage to a 50 mL conical tube.

Transfer ~40 mL to a different 50 mL as a retention sample and store the sample at ~4 °C.

Add 10.0 mL of ACN to the sample container, and vortex for 3 minutes.

Add 5 g QuEChERS salt packet to the sample, and manually shake/vortex for ~10 seconds, followed by opening the lid to de-gas.

Vortex the contents for 3 minutes followed by centrifuging for 5 min at 3000-4000 rpm.

Transfer 3.0 mL of the top layer to a 5 mL tube container with 0.75 mL of water.

Pipette 2.5 mL of the solution into 3 mL Captiva EMR Lipid cartridge on the positive pressure manifold with 15 mL Falcon collection tube.

Allow the liquid to drip by gravity until empty or use the pressure manifold to deliver a drop every 3-5 seconds on the low-pressure setting.

Add 1.5 mL of 80:20 ACN:H2O solution.

Allow the liquid to drip by gravity until empty or use the pressure manifold to deliver a drop every 3-5 seconds on the low-pressure setting.

Once there is no more standing liquid left in the EMR tube, increase the pressure to ~6 – 15 PSI on the low-pressure setting for about 30 seconds to elute all the remaining liquid from the tube.

Cap the 15 mL Falcon tube, briefly vortex, and vial (DF=2).

## Preparation of Edible Samples

Weigh 250.0 ± 5.0 mg of ground sample into a labeled, 50 mL tube.

Add 20.0 mL of H2O.

Vortex sample at speed 10 for 5 minutes, then sonicate for 5 minutes. Gentle heating while sonicating at ~35 °C may be helpful. Continue vortexing and sonicating until the sample is dissolved and/or evenly dispersed in the tube.

Add 10.0 mL of ACN to the sample container, and vortex for 1 minute.

Add all the contents of the 7.5g QuEChERS salt packet to the sample, and manually shake/vortex for 10 seconds, followed by opening the lid to de-gas.

Vortex the contents for 5 minutes.

Centrifuge sample at 3000-4000 rpm for 5 minutes.

Attach a clean, 0.2 µm PTFE filter to a syringe, dispense contents into a new 5 mL tube, and vial (DF=40).

Each cannabinoid will have its derived concentration multiplied by the correction factor listed in Table 10.

Table 10. Correction Factor in Edibles Matrices

| Analyte | Correction Factor |
| --- | --- |
| CBDV | 1.14 |
| CBDA | 1.08 |
| CBGA | 1.09 |
| CBG | 1.13 |
| CBD | 1.15 |
| THCV | 1.14 |
| CBN | 1.15 |
| d9-THC | 1.15 |
| d8-THC | 1.13 |
| CBC | 1.09 |
| THCA | 1.07 |

## Preparation of Transdermal Patch Samples

Cut out the patch and place into a 50 mL tube.

Add 20.0 mL of ACN to the sample container.

Vortex sample at speed 10 for 5 minutes, then sonicate for 5 minutes.

Vortex sample at speed 10 for 1 minute, then sonicate for 5 minutes.

If necessary, attach a clean, 0.2 µm PTFE filter to a syringe, and dispense contents into a new 5 mL tube (DF=100).

It is recommended to dilute the samples with ACN to 1000x.

## HPLC Setup and Measurement

See Table 11 for instrument parameters. Flow rate, oven temperature, injection volume (or the injection program), run time, and gradient should not be changed.

- Table 11. Instrument Parameters

| Parameters | Setting |
| --- | --- |
| Flow Rate | 1.4 mL/min |
| Oven Temperature | 45.0°C |
| Autosampler Temperature | 15.0°C |
| Injection Volume | 5 µL |
| Run Time | 9 min |
| Starting Wavelength | 190 nm |
| Ending Wavelength | 300 nm |
| PDA Cell Temperature | 40.0°C |

See Table 12 for gradient profile.

- Table 12. Gradient Profile

| Time (min) | %B |
| --- | --- |
| 0.00 | 70 |
| 8.00 | 80 |
| 8.01 | 70 |
| 9.00 | 70 |

See Table 13 for an example assay sequence template.

- Table 13. Example Assay Sequence Template

| Injection Order | Sample | Injection Type | Level | Number of Injections |
| --- | --- | --- | --- | --- |
| 1 | CCV | Control | 7 | ≥3 |
| 2 | Standards 1-6 (Optional if new curve is needed) | Standard | 1-6 | 1 |
| 3 | ACN Blank | Unknown | 0 | 1 |
| 4 | CCV | Control | 7 | 1 |
| 5 | LOQ | Control | 1 | 1 |
| 6 | LCS(s) | Control | 8, 9 (Beverage QC) | 1 |
| 7 | Sample(s) | Unknown | 0 | 1 |
| 8 | CCV | Control | 7 | 1 |

## Quality Control Requirements

All Quality Control Requirements must be performed and pass the applicable acceptance criteria, before any sample analysis. Any analytes that do not meet all criteria must be re-injected in the same or different sequence for those analytes that contain reportable results for any sample(s) assayed. Sequences may be stopped early, and restarted, to address any issues relating to Quality Control.

To demonstrate accuracy and precision, the CCV is injected a minimum of three times. The last three injections must have an accuracy of ±15% of the nominal concentration (20 ppm) for each cannabinoid. The precision of the last three injections must have an %RSD ≤5% (using sample standard deviation) for all cannabinoids.

The resolution between every cannabinoid in the CCV injections should be ≥1.0.

If a calibration curve is run, the back calculated concentration at each level must ±10% of the nominal concentration for each cannabinoid. The coefficient of determination (r2) must be ≥0.99 for all cannabinoids.

Inject ACN as a reagent blank. The blank injection must be free of any interfering peaks and/or background contamination.

The LOQ should have an accuracy of ±20% for each cannabinoid.

Laboratory Control Samples have an acceptance criterion of the ±20% nominal concentration for each cannabinoid. LCSs outside of the ±20% should be reprepared/extracted along with any concurrently prepared samples of the same matrix type if it is determined that the HPLC is not at fault. Only the non-carboxylic acid moieties of the cannabinoids in the scope of this SOP will be analyzed in Beverage and Edible Laboratory Control Samples.

One sample duplicate must be prepped every 20 samples. Duplicate samples can be from any of the matrices described above. Duplicate LCSs may be used as well. The reportable results (concentration) for each Duplicate must have a %Difference of ≤±20% (Equation 1) for each analyte.

- %Difference = Sample 1 - Sample 2Average(Sample 1 & Sample 2) × 100%  		(1)

- If samples are prepped in duplicate, then the reportable results of the first sample injected will be used.

CCV injections must bracket every 10 sample and LCS injections. The accuracy of each CCV injection must be ±15% of the nominal concentration for each cannabinoid.

## Calculations and Reporting of Results

Process the batch run results using quadratic regression with 1/C2 weighting at 220 nm in LabSolutions.

Manual peak identification and integration are permitted if it is apparent relevant peaks are not properly processed by the software.

A sample is positive for any of the cannabinoids in the scope of this method if the peaks in the sample correspond to the retention time of the peaks in the CCV injections and are above the LOQ. See Table 14.

- Table 14. Relative Retention Times

| Analyte | Relative RT |
| --- | --- |
| CBDV | 0.67 |
| CBDA | 0.85 |
| CBGA | 0.89 |
| CBG | 0.95 |
| CBD | 1.00 |
| THCV | 1.06 |
| THCVA | 1.26 |
| CBN | 1.35 |
| Δ9-THC | 1.61 |
| Δ8-THC | 1.66 |
| CBL | 1.83 |
| CBC | 1.90 |
| THCA | 1.96 |

Any unknown sample peaks greater than the lower limit of quantitation (LOQ), but less than the upper limit of quantitation (ULOQ; the highest standard point), must be reported.

Any unknown sample peaks greater than the ULOQ should be diluted with ACN and re-injected, such that they fall between the LOQ and ULOQ.

After processing the data in LabSolutions, export the batch data file as .CSV file and upload to QBench. All CoAs will be generated in QBench.

In QBench, Total THC will be reported for all samples using Equation 2.

- Total THC = ∆9-THC + (0.877 x THCA)		(2)

In QBench, Total CBD will be reported for all samples using Equation 3.

- Total CBD = CBD + (0.877 x CBDA)		(3)

Any cannabinoid not present will be reported as <LOQ.

Cannabinoid Purity will be assessed in QBench for lower-potency hemp edible products and hemp-derived consumer products. The Weight % of all impurities, quantitated using the CBD calibration curve, must be <5% compared to ∆9-THC.

For products containing artificially sourced cannabinoids, the concentration of cannabinoids CBDV, CBDA, CBGA, THCV, THCVA, Δ8-THC, CBL, and THCA will be added together and must be 0.5 mg/serving.

If cannabinoids are listed on sample labels, they must be within 15% of the label claim. For non-intoxicating cannabinoids present at low levels (< 5%) in products with more than 70% THC, the allowed variance is 25% of the label amount.

- .

# Change History

| Revision # | Change History |
| --- | --- |
| 1.0 | - New Document |
| 1.1 | - Section 8.7.4 the amount of acetonitrile was changed from 10 mL to 20 mL. - Section 8.7.8 the DF was updated from 20 to 40. - Edible correction factors were listed in Table 10 in Section 8.10.9. |