|  | Document Title: |  | Revision Date: | 3/30/2026 |
| --- | --- | --- | --- | --- |
|  | Method Transfer Report for Cannabinoid Potency Assay |  | Version Number: | 1.0 |
|  |  |  | Approved By: | Andrew Miller |
|  |  |  | Department/Category: | Laboratory |
|  | Document Type: | Report | Page(s): | 1 of 6 |

# Purpose

The purpose of this report is to summarize the results from the method transfer of the “Cannabinoid Potency Assay SOP” for its intended use as a potency assay to quantitatively determine the cannabinoid content in marijuana and hemp flower/plant products as well as marijuana/hemp derived products at the AIT lab in Moorhead, MN.

# Scope

The scope of this document is to present the results from the execution of the validation for “Cannabinoid Potency Assay SOP.” See below for the compounds to be validated.

          - Delta 9-tetrahydrocannabinol (d9-THC)

          - Delta 8-tetrahydrocannabinol (d8-THC)

          - Tetrahydrocannabinolic acid (THCA)

          - Tetrahydrocannabivarin (THCV)

          - Cannabidiol (CBD)

          - Cannabidiolic acid (CBDA)

          - Cannabidivarin (CBDV)

          - Cannabigerol (CBG)

          - Cannabigerolic acid (CBGA)

          - Cannabinol (CBN)

          - Cannabichromene (CBC)

# Summary

## Study Design

This validation report expands the scope of this method from its initial validation.

Per ICH Q2(R2) “Validation of Analytical Procedures”, the following parameters were addressed: Specificity, Accuracy, Precision – Repeatability for a potency assay (Table 1).

The matrices used in this validation were cellulose powder, olive oil, carbonated beverage, and gelatin/sugar powder.

## Results Summary

The results of the validation study are summarized in Table 1.

- Table 1. Validation Summary

| - Validation Parameter | - Acceptance Criteria | - Results | - Outcome |
| --- | --- | --- | --- |
| - Specificity | - Each analyte is present in each spiked matrix | - Each analyte was confirmed | - Pass |
| - Accuracy | - 80% – 120% mean Recovery of the nominal concentration for each analyte | - 80% – 120% | - Pass |
| - Precision – Repeatability | - RSD ≤20% for every analyte | - ≤20% for each analyte | - Pass |

## Study Activities

The validation was conducted in accordance with the “Method Validation Protocol for Cannabinoid Potency Assay” and “Cannabinoid Potency Transfer Protocol” in accordance to ICH Q2(R2).

The assays that were conducted to complete this validation: 20260318_CP and 20260319_CP_2

# Study Samples and Results

## Study Samples

Flower: Cellulose powder from MP Biomedicals was used as a surrogate for cannabis flower samples due to interferences of native cannabinoids.

Concentrate: Olive oil from Absolute Standards, Inc. was used as the matrix for cannabis concentrates.

Beverage: Mojito La Croix beverages were used as the matrix for cannabis infused beverages.

Edibles: Cherry gelatin dessert powder containing sugar was used as the matrix for edibles.

All samples were spiked, extracted, and analyzed per the “Cannabinoid Assay SOP” as well as the “Cannabinoid Potency Transfer Protocol” and document on the “Cannabinoid Assay Form” and “Cannabinoid Potency Validation Assay Form.”

# Results

## Specificity

Specificity was demonstrated with each analyte confirmed in every spiked sample.

## Accuracy

Experimental Design

Each matrix was spiked and extracted as described in the SOP.

Results

The mean %Accuracy (%Recovery) results are shown in Table 2. The mean back calculated concentrations for of each analyte were between 80% – 120%, thus demonstrating the accuracy of this method.

- Table 2.  %Recovery Results Extracted from each Matrix

| Analyte | Flower Rep 1 | Flower Rep 2 | Flower Rep 3 | Conc. Rep 1 | Conc. Rep 2 | Conc. Rep 3 | Bev. Rep 1 | Bev. Rep 2 | Bev. Rep 3 | Edible Rep 1 | Edible Rep 2 | Edible Rep 3 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| CBDV | 104.3 | 94.4 | 97.4 | 98.0 | 100.4 | 94.5 | 83.6 | 95.3 | 96.8 | 86.0 | 87.1 | 89.3 |
| CBDA | 98.5 | 98.5 | 98.0 | 102.0 | 103.0 | 101.3 | 108.5 | 106.8 | 110.1 | 93.2 | 91.9 | 93.4 |
| CBGA | 99.7 | 97.1 | 96.2 | 103.6 | 104.5 | 101.8 | 101.9 | 100.6 | 106.3 | 92.0 | 91.4 | 91.3 |
| CBG | 98.4 | 96.3 | 95.8 | 98.6 | 100.7 | 100.7 | 101.7 | 98.0 | 99.6 | 90.1 | 86.4 | 88.7 |
| CBD | 95.7 | 94.0 | 94.0 | 99.7 | 99.5 | 96.3 | 97.0 | 93.1 | 96.0 | 85.9 | 87.2 | 86.9 |
| THCV | 96.7 | 95.4 | 97.1 | 100.5 | 100.5 | 98.0 | 101.1 | 97.0 | 98.6 | 88.1 | 87.3 | 88.3 |
| CBN | 96.2 | 95.5 | 95.2 | 98.7 | 98.8 | 97.5 | 100.9 | 103.1 | 99.6 | 87.5 | 86.6 | 87.5 |
| d9-THC | 95.9 | 91.7 | 92.3 | 99.2 | 96.7 | 96.4 | 100.6 | 98.6 | 100.3 | 89.3 | 83.8 | 87.5 |
| d8-THC | 95.8 | 93.7 | 93.3 | 98.6 | 100.4 | 94.9 | 97.5 | 99.6 | 94.4 | 87.3 | 87.8 | 90.1 |
| CBC | 101.4 | 100.2 | 101.2 | 102.5 | 104.3 | 102.4 | 103.2 | 103.2 | 106.7 | 88.8 | 90.0 | 95.7 |
| THCA | 96.6 | 97.4 | 98.6 | 104.1 | 106.0 | 106.7 | 105.0 | 103.8 | 104.6 | 91.3 | 93.9 | 94.3 |

## Precision – Repeatability

Experimental Design

The beverage and olive oil matrices were spiked as described in the protocol.

Results

The %RSD results are shown in Table 3. Every analyte achieved a %RSD of ≤20%.

- Table 3.  %RSD Repeatability Results Extracted from each Matrix

| Analyte | Flower | Concentrate | Beverage | Edible |
| --- | --- | --- | --- | --- |
| CBDV | 5.2 | 3.0 | 7.8 | 1.9 |
| CBDA | 0.3 | 0.9 | 1.6 | 0.9 |
| CBGA | 1.9 | 1.3 | 2.9 | 0.4 |
| CBG | 1.4 | 1.2 | 1.8 | 2.1 |
| CBD | 1.1 | 2.0 | 2.1 | 0.8 |
| THCV | 1.0 | 1.5 | 2.1 | 0.6 |
| CBN | 0.6 | 0.7 | 1.7 | 0.6 |
| d9-THC | 2.5 | 1.6 | 1.1 | 3.3 |
| d8-THC | 1.4 | 2.9 | 2.7 | 1.7 |
| CBC | 0.6 | 1.0 | 1.9 | 4.0 |
| THCA | 1.0 | 1.3 | 0.6 | 1.7 |

# On-Going Method Verification

## Robustness

Experimental Design

On-Going Quantitative Method Verification should be assessed by examining the Spiked LCS from every assay.

Acceptance Criteria

The acceptance criteria for the %RSD is ≤20% and the %Recovery for is 80% – 120% for each analyte in the Spiked LCS samples.

## Measurement Uncertainty (MU)

Measurement uncertainty is calculated by trending the Robustness data of this method using each Spiked LCS for every analyte. Measurement uncertainty (%) is determined using Equation 1.1:

- MU =2xMean bias2+Stdev.P bias2+%RSD2 	           		  (1.1)

- Where the Stdev.P is the standard deviation of the population in Excel and the Mean bias is the mean relative bias as expressed in Equation 1.2:

- Mean bias=Measured Concentration-Nominal ConcentrationNominal Concentration×100%	              	  (1.2)

See Table 4 for the determined method uncertainty values for each analyte from this validation.

- Table 4.  Expanded Method Uncertainty% (MU%) from All LCS Spiked Solutions

| Analyte | Flower | Concentrate | Beverage | Edible |
| --- | --- | --- | --- | --- |
| CBDV | 13.5 | 9.3 | 25.4 | 25.5 |
| CBDA | 3.4 | 4.8 | 17.4 | 14.6 |
| CBGA | 6.7 | 7.5 | 9.5 | 16.9 |
| CBG | 7.3 | 3.1 | 4.8 | 23.8 |
| CBD | 11.2 | 5.8 | 10.7 | 26.7 |
| THCV | 7.6 | 3.9 | 5.8 | 24.2 |
| CBN | 8.9 | 3.9 | 5.0 | 25.6 |
| d9-THC | 14.7 | 6.5 | 2.8 | 27.5 |
| d8-THC | 12.0 | 8.4 | 8.9 | 23.6 |
| CBC | 2.5 | 6.7 | 10.1 | 19.8 |
| THCA | 5.6 | 11.7 | 9.1 | 14.3 |

# Nonconforming Events and Quality Events

Both flavors of the La Croix beverages contained an interfering peak with CBDA. The area of this interfering peak was below the Standard 1 area from the blank injections. However, this interfering peak was still integrated, and that concentration was subtracted from every beverage spiked CBDA concentration.

The chromatography of THCA contained a split peak in couple of injections from the Concentrate Day 1 injections. These samples were re-injected later in the same sequence, and that data was used towards this validation.

# Conclusions

- Acceptance criteria were met every analyte in the scope of this method transfer for Specificity, Accuracy, and Precision – Repeatability. Robustness is on-going with performance assessed from Spiked LCS samples. The analytes listed above are considered validated for all matrices in the SOP “Cannabinoid Potency Assay SOP.” Addendums to this validation are permitted for the expansion of the scope of this method.

# References

Table 6. Document References

Document Title

Cannabinoid Potency Assay SOP

Cannabinoid Potency Transfer Protocol

Cannabinoid Potency Assay Form

Cannabinoid Potency Validation Sample Preparation Form

ICH Q2(R2) Guideline on validation of analytical procedures

20260318_CP

20260319_CP_2

# Change History

| Revision # | Change History |
| --- | --- |
| 1.0 | New Document |